## Summary
¦ | Scenario | ¦ | Description | ¦ | Good | ¦ | Decent | ¦ | Bad | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Baroreceptors |  | Patient has 10% blood loss in 30s. View response in heart rate, contractility, systemic vascular resistance, and venous compliance. | ¦<span class="success"> | 2 | </span>¦<span class="warning"> | 1 | </span>¦<span class="danger"> | 0 | </span>¦
¦ | TBI |  | Patient has three brain injuries of increasing severity with recovery in between. View response in intracranial pressure, cerebral blood flow, and cerebral perfusion pressure. | ¦<span class="success"> | 19 | </span>¦<span class="warning"> | 6 | </span>¦<span class="danger"> | 0 | </span>¦
¦ |  | ¦ | Total | ¦<span class="success"> | 21 | </span>¦<span class="warning"> | 7 | </span>¦<span class="danger"> | 0 | </span>¦

## Key
¦ | Key | ¦
¦ | --- | ¦
¦<span class="success"> | Good agreement: correct trends or <10% deviation from expected | </span>¦
¦<span class="warning"> | Some deviation: correct trend and/or <30% deviation from expected | </span>¦
¦<span class="danger"> | Poor agreement: incorrect trends or >30% deviation from expected | </span>¦

## Baroreceptors
Baroreceptors
¦ | Action | ¦ | Notes | ¦ | Sampled Scenario Time (s) | ¦ | Heart Rate (beats/min) | ¦ | Cardiac Output (mL/min) | ¦ | Systemic Vascular Resistance (mmHg s/mL) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Hemorrhage | ¦ | 10% blood loss in 30 s (Class 1 Hemorrhage) | ¦ | 325 | ¦<span class="success"> | Slight increase, remain < 100 @cite gutierrez2004clinical | </span>¦<span class="warning"> | Decrease ~15-20% @cite Hosomi1979effect @cite Ottesen2004applied | </span>¦<span class="success"> | Increase 10-15% @cite Hosomi1979effect @cite Ottesen2004applied | </span>¦

## Sleep
Sleep
¦ | Action | ¦ | Notes | ¦ | Sampled Scenario Time (min) | ¦ | Glucose clearance (%/min) | ¦ | Glucose effectiveness (%/min) clearance without insulin | ¦ | Insulin response (pmol/min) | ¦ | Stanford sleepyness score (unitless) | ¦ | Cerebral glucose uptake (% from baseline) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Sleep of 4hr per day for 5 days | ¦ | sleep deprivation study taken from spiegle, given carb rich meal (62%) every 5h | ¦ | 7200 | ¦<span class="success"> | 40% lower (1.42%/min compared to 2.4%/min @cite spiegel1999impact | </span>¦<span class="warning"> | Decrease 30% @cite spiegel1999impact | </span>¦<span class="success"> | Decrease from 432 to 304 pmol/min @cite spiegel1999impact | </span>¦<span class="success"> | Increase from 2.1 to 4.4 @cite spiegel1999impact | </span>¦<span class="success"> | Decrease of 7% @cite spiegel1999impact | </span>¦

## TBI
TBI
¦ | Action | ¦ | Notes | ¦ | Action Occurrence Time (s) | ¦ | Sampled Scenario Time (s) | ¦ | Intracranial Pressure (mmHg) | ¦ | Cerebral Blood Flow (mL/min) | ¦ | Cerebral Perfusion Pressure (mmHg) | ¦ | Heart Rate (1/min) | ¦ | Respiration Rate (1/min) | ¦
¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦ | --- | ¦
¦ | Brain Injury | ¦ | Severity 0.2, CPP=MAP-ICP, StandardMale brain mass=1450g | ¦ | 20 | ¦ | 600 | ¦<span class="warning"> | 10% Increase @cite bergeronSME | </span>¦<span class="success"> | Decrease @cite steiner2006monitoring | </span>¦<span class="success"> | Decrease @cite balestreri2006impact | </span>¦<span class="warning"> | 0-10% Decrease @cite bergeronSME | </span>¦<span class="warning"> | 0-10% Decrease @cite bergeronSME | </span>¦
¦ | Brain Injury | ¦ | Severity 0 | ¦ | 620 | ¦ | 900 | ¦<span class="success"> | 7-15 mmHg @cite steiner2006monitoring | </span>¦<span class="success"> | 50-65 mL/100g/min @cite guyton2006medical | </span>¦<span class="success"> | 60-98 mmHg | </span>¦<span class="success"> | 72 @cite guyton2006medical | </span>¦<span class="success"> | [12.0, 20.0], [13.0, 19.0] @cite silverthorn2013human @cite mantoni2007reduced | </span>¦
¦ | Brain Injury | ¦ | Severity 0.75 | ¦ | 920 | ¦ | 1700 | ¦<span class="success"> | Increase @cite steiner2006monitoring | </span>¦<span class="success"> | Decrease @cite steiner2006monitoring | </span>¦<span class="success"> | Decrease @cite balestreri2006impact | </span>¦<span class="warning"> | 0-15% Decrease @cite bergeronSME | </span>¦<span class="warning"> | 0-20% Decrease @cite bergeronSME | </span>¦
¦ | Brain Injury | ¦ | Severity 0 | ¦ | 1720 | ¦ | 2000 | ¦<span class="success"> | 7-15 mmHg @cite steiner2006monitoring | </span>¦<span class="success"> | 50-65 mL/100g/min @cite guyton2006medical | </span>¦<span class="success"> | 60-98 mmHg | </span>¦<span class="success"> | 72 @cite guyton2006medical | </span>¦<span class="success"> | [12.0, 20.0], [13.0, 19.0] @cite silverthorn2013human @cite mantoni2007reduced | </span>¦
¦ | Brain Injury | ¦ | Severity 1 | ¦ | 2020 | ¦ | 3220 | ¦<span class="success"> | >25 mmHg @cite steiner2006monitoring | </span>¦<span class="success"> | <8 mL/100g/min @cite steiner2006monitoring | </span>¦<span class="success"> | Decrease @cite balestreri2006impact | </span>¦<span class="warning"> | Decrease @cite bergeronSME | </span>¦<span class="warning"> | Decrease @cite bergeronSME | </span>¦
Note: In BG, this is Brain-Inflow